SERVICE & MAINTENANCE OF EQUIPMENT

Some equipment in Certified Facilities requires mandatory maintenance and certification. These may be dealt with centrally by your Building Manager so check with them first.

Autoclaves: –

If used for infectious waste and/or GM material, a record of service, calibration and certification is required at least annually.

You should keep here a copy of the latest certificate here.

Monthly verification (e.g. spore testing or biological indicator testing) must be conducted monthly in-house and records may be kept in a separate folder.

Please note here where the annual and monthly records are kept:

_________________________________________________________________________________

Approved cycles for infectious & GM Waste:

It is recommended that you autoclave at 121°C for 20 minutes minimum.

AS2243.3 state the following: 15 min at 121°C and 103KPa Or 3 min at 134°C and 203KPa

Note: 	The densest & deepest part of the load must be exposed to these minimum requirements, you may need to run your autoclave for longer if you have very large loads, or large volumes of liquid or dense items such as soil and carcasses.

Approved Tests:

Biological indicators –  which require incubation e.g. 3M Attest & EZ-Test

Bacterial enzyme indicators – e.g. EZ-Test

Chemical indicators - that progressively change colour with the time exposed at specified temperatures. A class 5 or 6 chemical indicator should be used. e.g. TITEMS Diagnostic emulator (class 6 indicator) or 3M Comply Steam Chemical Integrator (class 5 indicator).

Note: 	The indicator must be placed in the densest and deepest part of the load.

Note:-	If the autoclave fails the test, you must stop using it to process waste. You must redo the test. If the test fails a second time, you must place a notice on the autoclave that it is ‘out of service’ then you must arrange for the autoclave to be serviced and fixed.

_________________________________________________________________________________

Biosafety Cabinets:-

Must be serviced and certified at least annually.

You should keep here a copy of the latest certificate.

_________________________________________________________________________________

Centrifuges:-

If used for GM material, must have an SOP for dealing with spills,

It is recommended you keep a copy of that SOP here.

Document any spills and cleanup in the table below. Note below any regular maintenance or servicing of the centrifuge used for GM.

SERVICE & MAINTENANCE OF EQUIPMENT Con’t…

| DATE | EQUIPMENT | SERVICE/MAINTENANCE | COMPANY |
| --- | --- | --- | --- |
| xx/xx/xx | Autoclave – (name/serial/location) | Annual calibration Equipment failure service | Atherton/Getinge |
| xx/xx/xx | Biosafety Cabinet (name/serial/location) | Annual service | Camphil Far |

Note that you may keep this information in your own database or folder, but record the location here as it must be accessible to the Biosafety Advisors.